Supplemental Materials

Home-based Video-Polysomnography for Sleep-Related Motor Behaviors: Development, Feasibility, and Diagnostic Performance at the Bologna Sleep Center

Greta Mainieri1, Luca Baldelli1,2*, Francesco Mignani1, Francesca Cotichelli2, Giuseppe Loddo3, Filomena Miele1,2, Ezechiele Foschini1,2, Angelica Montini2, Felice Di Laudo2, Caterina Pazzaglia1, Monica Sala1, Federica Provini1,2

1 IRCCS Istituto delle Scienze Neurologiche di Bologna, Bologna, Italia

2 Department of Biomedical and NeuroMotor Sciences, University of Bologna, Bologna, Italy

3 Department of Primary Care, Azienda AUSL di Bologna, Bologna, Italy

*Corresponding author

Luca Baldelli

IRCCS, Institute of Neurological Sciences of Bologna

Bellaria Hospital, Via Altura, 3 40139 Bologna (ITALY)

E-mail: luca.baldelli4@unibo.it

Table S1. Technical problems in NREM parasomnia/epilepsy and RBD suspicion groups

|  | Overall technical problems (n=51) | NREM parasomnia/epilepsy suspicion | RBD suspicion |
| --- | --- | --- | --- |
| Exam related | 16 | 9 | 7 |
| Interruption of recording | 5 (31.25%) | 3 (33.33%) | 2 (28.57%) |
| Desynchronization video-EEG | 5 (31.25%) | 4 (44.45%) | 1 (14.29%) |
| Polygraphic artefacts | 3 (18.75%) | 0 (00.00%) | 3 (42.85%) |
| Not downloadable recording | 2 (12.50%) | 1 (11.11%) | 1 (14.29%) |
| Spotlight malfunctioning | 1 (6.25%) | 1 (11.11%) | 0 (0.00%) |
| Patient related | 35 | 24 | 11 |
| Video not switched on (totality or a part of the recording) | 13 (37.14%) | 7 (29.17%) | 6 (54.55%) |
| Unchanged batteries | 10 (28.57%) | 9 (37.50%) | 1 (9.09%) |
| Infrared light not switched on | 9 (25.72%) | 6 (25.00%) | 3 (27.27%) |
| Camera not connected to power supply | 2 (5.71%) | 2 (8.33%) | 0 (00.00%) |
| Incorrect positioning of camera | 1 (2.86%) | 0 (0.00%) | 1 (9.09%) |

Qualitative variables are expressed as absolute number and percentage.

Figure S1. Technical artefacts in the first 50 consecutive Home VPSGs

The top panel displays the percentage of channel artefacts in the first 50 consecutive home VPSGs. Polygraphic recordings with less than 10% channel artefacts are shown in green, while those with a mean artefact percentage between 10% and 20% are shown in yellow. Exams with an artefact percentage exceeding 20% are indicated in red.

In the bottom panel, numbers above the red line represent the mean artefact percentage over time.