Table S5: Detailed information on the medical physics (MP) answers about their equipment in the questionnaire (22 departments). Abbreviations: CT: Computed Tomography, n/s: not specified; Linac: linear accelerator; MR: magnetic resonance.

| MP (22) | Equipment | Planning CT |  | Slice thickness |  | Consistency test |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 16 slices | 7 (32%) | 1 mm | 6 (27%) | yes | 22 (100%) |
|  |  | 64 slices | 6 (27%) | 1.5 mm | 3 (14%) | due to national or | 18 (82%) |
|  |  | other | 4 (18%) | 1-2 mm | 7 (32%) | international recommendations | 7 (32%) |
|  |  | n/s | 5 (23%) | other | 6 (27%) | n/s | 1 (5%) |
|  |  | Dose calculation algorithm |  | Grid Size |  |  |  |
|  |  | Type A | 3 (14%) | 1 mm | 3 (14%) |  |  |
|  |  | Type B | 11 (50%) | 1-2 mm | 8 (36%) |  |  |
|  |  | Monte Carlo | 11 (50%) | 2 mm | 10 (45%) |  |  |
|  |  |  |  | other | 1 (5%) |  |  |
|  |  | Treatment machine |  | Audit |  | Intervall |  |
|  |  | Conv. Linac | 15 (68%) | yes | 13 (59%) | yearly | 6 (46%) |
|  |  | CyberKnife | 5 (23%) | no | 7 (32%) | every two years | 3 (23%) |
|  |  | MR-Linac | 3 (14%) | n/s | 2 (9%) | other | 2 (15%) |
|  |  | Synchrotron | 1 (5%) |  |  | n/s | 2 (15%) |
|  |  | n/s | 2 (9%) |  |  |  |  |

Table S6: Detailed information on the medical physics (MP) answers about quality assurance in the questionnaire (22 departments). Abbreviations: n/s: not specified; IROC: Imaging and Radiation Oncology Core; SBRT: stereotactic body radiotherapy, IMT: intensity modulated technique, QA: quality assurance.

| MP (22) | Quality Assurance | National guidelines small field |  | Regular Dose audits |  | Daily control isocenter consistency (Winston Lutz) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | yes | 2/8 (25%) | yes | 16 (73%) | yes |  | 16 (73%) |  |  |
|  |  | no | 2/8 (25%) | no | 5 (23%) | no |  | 4 (18%) |  |  |
|  |  | different answers by same country | 4/8 (50%) | n/s | 1 (4%) | n/s |  | 2 (9%) |  |  |
|  |  |  |  | Who performs the audit? |  |  |  |  |  |  |
|  |  |  |  | External: 15 (94%), national or IROC Houston |  |  |  |  |  |  |
|  |  | End-to-End specifically for SBRT |  | Phantom type |  | Evaluation criteria |  | Intervall |  |  |
|  |  | yes | 17 (77%) | moving | 5 (29%) | 1 mm: 10 (59%) | max. 2%: 2 (12%) | monthly |  | 6 (35%) |
|  |  | no | 5 (23%) | non-moving | 6 (35%) | 2 mm: 3 (18%) | max. 3%: 7 (41%) | quarterly |  | 3 (18%) |
|  |  |  |  | both | 4 (24%) | other: 3 (18%) | up to 5%: 6 (35%) | yearly |  | 5 (29%) |
|  |  |  |  | n/s | 2 (12%) | n/s: 1 (6%) | n/s: 2 (12%) | other |  | 3 (18%) |
|  |  | Patient-specific plan verification |  | Evaluation criteria |  |  |  |  |  |  |
|  |  | yes, for all | 8 (36%) | Point dose | 10 (46%) | Gamma 3%/1mm |  | 3 (15%) |  |  |
|  |  | yes, only IMT | 10 (46%) | Gamma 3%/3mm | 5 (25%) | Gamma 5%/1mm |  | 2 (10%) |  |  |
|  |  | yes, infrequent | 4 (18%) | Gamma 2%/2mm | 5 (25%) | other |  | 5 (25%) |  |  |
|  |  | Regular accuracy check, according to |  | Inspection of QA |  | Offline review |  | Post-treatment analysis |  |  |
|  |  | system specific | 12 (55%) | yes | 17 (77%) | yes, regular | 7 (32%) | yes | 8 (36%) |  |
|  |  | own defined | 8 (36%) | internal | 1 (6%) | yes, infrequent | 6 (27%) | no | 12 (55%) |  |
|  |  | national guidelines | 15 (68%) | independent | 15 (88%) | no | 7 (32%) | n/s | 2 (9%) |  |
|  |  | international guidelines | 10 (45%) | n/s | 1 (6%) | n/s | 2 (9%) |  |  |  |
|  |  | other | 1 (5%) | No | 5 (23%) |  |  |  |  |  |