Supplementary material to: External hair contamination from cannabis and “light cannabis” delivered by smoking and vaping: an in-vitro study

Arianna Giorgetti1, Paolo Fais1, Jennifer Paola Pascali1*, Susan Mohamed1, Francesca Rossi1, Marco Garagnani1, Guido Pelletti1

1Department of Medical and Surgical Sciences, Unit of Legal Medicine, University of Bologna, Via Irnerio 49, 40126 Bologna, Italy

* Corresponding author: jennifer.pascali@gmail.com

Content of the supplementary:

- Table A. Quantifier and qualifier ions, cone voltage and collision energy of the analytes of interest and of the ISs.

- Table B. Validation parameters, of the LC-MS/MS validated method for the quantification of Δ9-THC and CBD, showing linearity, limit of detection (LOD), lower limit of quantification (LLOQ), precision and accuracy.

- Figure A. Statistical comparison (paired parametrical t-test) of Δ9-THC content in unwashed and washed samples.

Table A. Quantifier and qualifier ions, cone voltage and collision energy of the analytes of interest and of the ISs.

| Analyte | Quantifier and qualifier ions | Cone voltage [V] | Collision energy [V] |
| --- | --- | --- | --- |
| Δ9-THC | 315.21→ 193.15 315.21→ 123.00 315.21→ 259.30 | 35 35 35 | 18 34 22 |
| Δ9-THC-D3 | 318.20→ 123.00 318.20→ 196.10 | 35 35 | 34 22 |
| CBD | 315.15→ 193.15 315.15→ 123.10 | 35 35 | 20 32 |
| CBD-D3 | 318.30→ 196.20 318.30→ 123.20 | 40 40 | 22 32 |

Table B. Validation parameters, of the LC-MS/MS validated method for the quantification of Δ9-THC and CBD, showing linearity, limit of detection (LOD), lower limit of quantification (LLOQ), precision and accuracy.

| Analytes of interest |  | Weighing | Calibration range 1 (n. calibration point) and R2 |  | Calibration range 2 (n. calibration point) and R2 |  |
| --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC |  | 1/x | 0.025-0.5 (4-points calibration curve) R2= 0.997 |  | 1-300 (6-points calibration curve) R2= 0.991 |  |
| CBD |  | 1/x | 0.1-2 (4-points calibration curve) R2= 0.994 |  | 1-300 (6-points calibration curve) R2= 0.992 |  |
|  | LOD [ng/mg] | LLOQ [ng/mg] | Intraday (n=5) |  | Interday (n=9) |  |
|  |  |  | Precision (CV%) | Accuracy (bias%) | Precision (CV%) | Accuracy (bias%) |
| Δ9-THC | 0.009 | 0.025 | 14.5 | - 6.3 | 11.0 | 7.4 |
| CBD | 0.03 | 0.1 | 5.6 | -3.8 | 1.6 | - 0.9 |

Linearity as tested by five calibration curves performed in five different days.

Precision (coefficient of variation or CV%) = 100 x standard deviation/mean.

Accuracy (bias%) = 100 x [(mean concentration found – nominal concentration)/nominal concentration].

LOD = defined by signal to noise 1:3 with respect to LLOQ.

Figure A. Statistical comparison (paired non-parametrical t-test) of Δ9-THC content in unwashed and washed samples. p-values < 0.05 for significance.