Supplementary material - Tables

Table S1: Generated experimental design

| t1 (min) | solv | ins | t2 (min) |
| --- | --- | --- | --- |
| 5 | MeOH | Ultrasound sonicator | 5 |
| 5 | ACN + 0,1% FA | Ultrasound sonicator | 5 |
| 5 | MeOH +EtOH | Ultrasound sonicator | 5 |
| 5 | MeOH + 0,1%FA | Ultrasound sonicator | 5 |
| 5 | ACN | Ultrasound sonicator | 5 |
| 5 | (MeOH + EtOH 1:1) + 0,1% FA | Ultrasound sonicator | 30 |
| 5 | ACN | Ultrasound sonicator | 30 |
| 30 | ACN + 0,1% FA | Ultrasound sonicator | 30 |
| 30 | MeOH + 0,1%FA | Ultrasound sonicator | 30 |
| 30 | MeOH +EtOH | Ultrasound sonicator | 30 |
| 30 | ACN | Ultrasound sonicator | 30 |
| 30 | MeOH | Ultrasound sonicator | 30 |
| 30 | (MeOH + EtOH 1:1) + 0,1% FA | Ultrasound sonicator | 5 |
| 30 | ACN | Ultrasound sonicator | 5 |
| 5 | (MeOH + EtOH 1:1) + 0,1% FA | Vortex | 5 |
| 5 | MeOH | Vortex | 5 |
| 5 | MeOH +EtOH | Vortex | 30 |
| 5 | ACN | Vortex | 30 |
| 5 | ACN + 0,1% FA | Vortex | 30 |
| 5 | MeOH + 0,1%FA | Vortex | 30 |
| 30 | MeOH + 0,1%FA | Vortex | 5 |
| 30 | MeOH +EtOH | Vortex | 5 |
| 30 | ACN | Vortex | 5 |
| 30 | ACN + 0,1% FA | Vortex | 5 |
| 30 | MeOH | Vortex | 30 |
| 30 | (MeOH + EtOH 1:1) + 0,1% FA | Vortex | 30 |

Table S2: MS/MS Parameters and HPLC Retention Times

| ASMs | Parent Ion (m/z) | Product Ion (m/z) | Polarity | Retention Time (min) |
| --- | --- | --- | --- | --- |
| Carbamazepine (CBZ) | 238,1 | 193,06 | positive | 3,71 |
| Lacosamide (LCM) | 252,1 | 108 | positive | 2,87 |
| Levetiracetam (LEV) | 172,1 | 70 | positive | 1,92 |
| Lamotrigine (LTG) | 256,1 | 210,9 | positive | 2,99 |
| Valproic acid (VPA) | 143.1 | 99.05 | negative | 4,26 |
| Internal standard |  |  |  |  |
| Carbamazepine-D10 | 247,1 | 204,05 | positive | 3,68 |
| Lacosamide-13C D3 | 254,1 | 108,05 | positive | 2,81 |
| Levetiracetam-D6 | 174,1 | 69,05 | positive | 1,99 |
| Lamotrigine-13C, 15N4 | 261,1 | 169,9 | positive | 1,95 |
| Valproic Acid-D6 | 149.1 | 106.05 | negative | 4,50 |

Table S3: Estimated coefficients, 95% confidence intervals and p values obtained from the fit of the model to the experimental pilot data. Significance levels of p values are marked with *,  ** and ***.

| Factors | Estimate coefficient | p value |
| --- | --- | --- |
| Intercept | 403± 22 | 2.8e-09*** |
| Time 1 | -0.99± 0.8 | 0.051 |
| Solv [MeOH] | -27±34 | 0.15 |
| Solv [MeOH + 0.1% FA] | -127±32 | 9.1e-05*** |
| Solv [MeOH + EtOH] | -33±32 | 0.074 |
| Solv [(MeOH + EtOH) + 0.1% FA] | -106±32 | 2.8e-04 *** |
| Solv [ACN + 0.1% FA] | 14±32 | 0.40 |
| Instrument [Vortex] | -25± 21 | 0.052 |
| Time 2 | -0.5 ± 0.4 | 0.058 |
| Time1: Solv [MeOH] | 0.19± 1.4 | 0.80 |
| Time1: Solv [MeOH + 0.1% FA] | 0.35± 1.3 | 0.61 |
| Time1: Solv [MeOH + EtOH] | 0.49± 1.3 | 0.48 |
| Time1: Solv [(MeOH + EtOH) + 0.1% FA] | 1.4± 1.3 | 0.079 |
| Time1: Solv [ACN + 0.1% FA] | -0.8± 1.3 | 0.25 |
| Solv [MeOH]: Instrument [Vortex] | 7.3± 33 | 0.68 |
| Solv [MeOH + 0.1% FA]: Instrument [Vortex] | 65± 33 | 0.0062** |
| Solv [MeOH + EtOH]: Instrument [Vortex] | 1.3± 33 | 0.94 |
| Solv [(MeOH + EtOH) + 0.1% FA]: Instrument [Vortex] | 18.8± 33 | 0.30 |
| Solv [ACN + 0.1% FA]: Instrument [Vortex] | -8.8± 33 | 0.62 |

Table S4: Linearity, intra-day, and inter-day precision and accuracy

| Intraday (n=6; 1 day) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Linearity |  |  |  |  |  |  |
| ASMs | R2 |  | Nominal Conc. (μg/mL) | Mean Conc. (μg/mL) | Std Dev. | %CV | %RE |
| CBZ | 0.998 |  |  |  |  |  |  |
|  |  | LLOQ | 0,4 | 0,38 | 0,04 | 10,53 | -5,26 |
|  |  | LowQC | 2 | 2,18 | 0,15 | 6,88 | 8,26 |
|  |  | MediumQC | 10 | 9,63 | 0,38 | 3,95 | -3,84 |
|  |  | HighQC | 15 | 14,5 | 0,47 | 3,24 | -3,45 |
| LCM | 0.998 |  |  |  |  |  |  |
|  |  | LLOQ | 0,5 | 0,48 | 0,009 | 1,88 | -4,17 |
|  |  | LowQC | 1 | 0,95 | 0,02 | 2,11 | -5,26 |
|  |  | MediumQC | 5 | 5,3 | 0,13 | 2,45 | 5,66 |
|  |  | HighQC | 10 | 9,48 | 0,45 | 4,75 | -5,49 |
| LTG | 0.999 |  |  |  |  |  |  |
|  |  | LLOQ | 0,5 | 0,46 | 0,006 | 1,30 | -8,70 |
|  |  | LowQC | 2 | 1,91 | 0,05 | 2,62 | -4,71 |
|  |  | MediumQC | 10 | 9,77 | 0,37 | 3,79 | -2,35 |
|  |  | HighQC | 20 | 19,06 | 1 | 5,25 | -4,93 |
| LEV | 0.998 |  |  |  |  |  |  |
|  |  | LLOQ | 1 | 0,94 | 0,13 | 13,83 | -6,38 |
|  |  | LowQC | 5 | 5,8 | 0,41 | 7,07 | 13,79 |
|  |  | MediumQC | 50 | 49,79 | 1,26 | 2,53 | -0,42 |
|  |  | HighQC | 80 | 91,43 | 1,53 | 1,67 | 12,50 |
| VPA | 0.996 |  |  |  |  |  |  |
|  |  | LLOQ | 10 | 9,67 | 3,17 | 32,78 | -3,41 |
|  |  | LowQC | 30 | 25,24 | 3,42 | 13,55 | -18,86 |
|  |  | MediumQC | 80 | 64,86 | 5,87 | 9,05 | -23,34 |
|  |  | HighQC | 120 | 104,33 | 7,8 | 7,48 | -15,02 |

| Interday (n=9; 3 days) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Linearity |  |  |  |  |  |  |
| ASMs | R2 |  | Nominal Conc. (μg/mL) | Mean conc. (μg/mL) | Std dev. | %CV | %RE |
| CBZ | 0.998 |  |  |  |  |  |  |
|  |  | LowQC | 2 | 2,14 | 0,04 | 1,87 | 6,54 |
|  |  | MediumQC | 10 | 9,67 | 0,37 | 3,87 | -3,41 |
|  |  | HighQC | 15 | 13,91 | 0,32 | 2,28 | -7,84 |
| LCM | 0.997 |  |  |  |  |  |  |
|  |  | LowQC | 1 | 0,92 | 0,01 | 1,30 | -8,70 |
|  |  | MediumQC | 5 | 5,47 | 0,09 | 1,61 | 8,59 |
|  |  | HighQC | 10 | 9,50 | 0,06 | 0,65 | -5,26 |
| LTG | 0.998 |  |  |  |  |  |  |
|  |  | LowQC | 2 | 1,76 | 0,12 | 6,70 | -13,64 |
|  |  | MediumQC | 10 | 10,80 | 0,23 | 2,15 | 7,41 |
|  |  | HighQC | 20 | 20,72 | 1,38 | 6,66 | 3,47 |
| LEV | 0.998 |  |  |  |  |  |  |
|  |  | LowQC | 5 | 5,62 | 0,14 | 2,42 | 11,03 |
|  |  | MediumQC | 50 | 55,23 | 0,03 | 0,05 | 9,47 |
|  |  | HighQC | 80 | 85,28 | 0,32 | 0,38 | 6,19 |
| VPA | 0.996 |  |  |  |  |  |  |
|  |  | LowQC | 30 | 22,46 | 6,00 | 26,71 | -33,57 |
|  |  | MediumQC | 80 | 102,10 | 0,54 | 0,53 | 21,65 |
|  |  | HighQC | 120 | 110,67 | 1,27 | 1,15 | -8,43 |

Table S5: Recovery (R%) and Matrix Effect (MF%) at different concentrations for ASMs

| RECOVERY (R%; n=3) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Conc. (μg/mL) |  | ASMs |  |
|  |  |  |  | CBZ |  |
| LOQ |  | 0,4 |  | 73 |  |
| LowQC |  | 2 |  | 83,2 |  |
| MediumQC |  | 10 |  | 74,5 |  |
| HighQC |  | 15 |  | 79 |  |
|  |  |  |  | LCM |  |
| LOQ |  | 0,5 |  | 74 |  |
| LowQC |  | 1 |  | 76 |  |
| MediumQC |  | 5 |  | 86 |  |
| HighQC |  | 10 |  | 83 |  |
|  |  |  |  | LTG |  |
| LOQ |  | 0,5 |  | 89 |  |
| LowQC |  | 2 |  | 98 |  |
| MediumQC |  | 10 |  | 86 |  |
| HighQC |  | 20 |  | 97 |  |
|  |  |  |  | LEV |  |
| LOQ |  | 1 |  | 108 |  |
| LowQC |  | 5 |  | 102 |  |
| MediumQC |  | 50 |  | 99 |  |
| HighQC |  | 80 |  | 104 |  |
|  |  |  |  | VPA |  |
| LOQ |  | 10 |  | 23 |  |
| LowQC |  | 30 |  | 21 |  |
| MediumQC |  | 80 |  | 44 |  |
| HighQC |  | 120 |  | 31 |  |
| MATRIX EFFECT at MEDIUM concentration (MF%; n=3) |  |  |  |
|  |  | ASMs |  |
|  |  | CBZ |  |
| Matrix 1 |  | 85,8 |  |
| Matrix 2 |  | 76,7 |  |
| Matrix 3 |  | 84,9 |  |
| Matrix 4 |  | 82,3 |  |
| Matrix 5 |  | 79,3 |  |
| Matrix 6 |  | 79,8 |  |
|  |  | LCM |  |
| Matrix 1 |  | 77,00 |  |
| Matrix 2 |  | 85,83 |  |
| Matrix 3 |  | 77,72 |  |
| Matrix 4 |  | 84,37 |  |
| Matrix 5 |  | 80,33 |  |
| Matrix 6 |  | 77,84 |  |
|  |  | LTG |  |
| Matrix 1 |  | 97,7 |  |
| Matrix 2 |  | 95,4 |  |
| Matrix 3 |  | 98,4 |  |
| Matrix 4 |  | 98,3 |  |
| Matrix 5 |  | 103,5 |  |
| Matrix 6 |  | 94,5 |  |
|  |  | LEV |  |
| Matrix 1 |  | 101,8 |  |
| Matrix 2 |  | 103,5 |  |
| Matrix 3 |  | 93,4 |  |
| Matrix 4 |  | 110,8 |  |
| Matrix 5 |  | 98,3 |  |
| Matrix 6 |  | 105,8 |  |
|  |  | VPA |  |
| Matrix 1 |  | 28,9 |  |
| Matrix 2 |  | 27,4 |  |
| Matrix 3 |  | 25,2 |  |
| Matrix 4 |  | 26,9 |  |
| Matrix 5 |  | 27,9 |  |
| Matrix 6 |  | 32,9 |  |

Table S6: LOD and LLOQ

| ASM | LLOD (ng/mL) | LLOQ (ng/mL) |
| --- | --- | --- |
| CBZ | 0,38 | 0,50 |
| LCM | 0,48 | 0,60 |
| LTG | 0,46 | 0,60 |
| LEV | 0,94 | 1,20 |
| VPA | 9,67 | 120 |

Table S7: Stability test

|  |  |  |  | 7 days stability at RT (N=3) |  |  |  |  |  | 7 days stability at -20 °C (N=3) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Nominal conc. (μg/mL) |  | Average | Std. Dev. | %CV |  | %RE |  | Average | Std. Dev. | %CV | %RE |
| CBZ | LowQC | 2 |  | 1,99 | -0,024 | -1,22 | -0,60 |  |  | 2,21 | 0,080 | 3,62 | 10,50 |
|  | MediumQC | 10 |  | 10,80 | 0,236 | 2,19 | 8,00 |  |  | 10,11 | 0,370 | 3,66 | 1,10 |
|  | HighQC | 15 |  | 13,89 | -0,899 | -6,47 | -7,39 |  |  | 16,95 | 0,250 | 1,47 | 13,00 |
| LCM | LowQC | 1 |  | 0,970 | 0,070 | 7,22 | -3,00 |  |  | 1,03 | 0,106 | 10,29 | 3,00 |
|  | MediumQC | 5 |  | 5,04 | 0,530 | 10,53 | 0,70 |  |  | 5,35 | 0,632 | 11,81 | 7,00 |
|  | HighQC | 10 |  | 9,94 | 0,279 | 2,80 | -0,62 |  |  | 10,6 | 1,200 | 11,32 | 6,00 |
| LTG | LowQC | 2 |  | 2,01 | 0,028 | 1,37 | 0,32 |  |  | 2,04 | 0,030 | 1,47 | 2,00 |
|  | MediumQC | 10 |  | 9,72 | 0,173 | 1,78 | -2,82 |  |  | 10,38 | 0,170 | 1,64 | 3,80 |
|  | HighQC | 20 |  | 20,50 | 2,723 | 13,28 | 2,51 |  |  | 21,17 | 0,470 | 2,22 | 5,85 |
| LEV | LowQC | 5 |  | 5,02 | 0,397 | 7,91 | 0,48 |  |  | 5,14 | 0,106 | 2,06 | 2,80 |
|  | MediumQC | 50 |  | 43,91 | 0,604 | 1,38 | -12,17 |  |  | 53,49 | 4,580 | 8,56 | 6,98 |
|  | HighQC | 80 |  | 89,35 | 9,323 | 10,43 | 11,69 |  |  | 90,14 | 4,100 | 4,55 | 12,68 |

|  |  |  |  | 15 days stability at RT (n=3) |  |  |  |  |  | 30 days stability at RT (n=3) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Nominal conc. (μg/mL) |  | Average | Std. Dev. | %CV |  | %RE |  | Average | Std. Dev. | %CV | %RE |
| CBZ | LowQC | 2 |  | 1,97 | -0,057 | -2,91 | -1,53 |  |  | 1,99 | 0,052 | 2,59 | -0,38 |
|  | MediumQC | 10 |  | 10,00 | 0,322 | 3,22 | 0,05 |  |  | 10,47 | 0,099 | 0,95 | 4,69 |
|  | HighQC | 15 |  | 15,01 | 0,004 | 0,02 | 0,04 |  |  | 14,35 | -0,001 | 0,00 | -4,33 |
| LCM | LowQC | 1 |  | 1,019 | 0,010 | 0,97 | 1,90 |  |  | 1,00 | 0,009 | 0,92 | -0,35 |
|  | MediumQC | 5 |  | 4,94 | -0,008 | -0,17 | -1,12 |  |  | 5,15 | 0,081 | 1,56 | 3,08 |
|  | HighQC | 10 |  | 10,16 | 0,127 | 1,25 | 1,60 |  |  | 9,73 | -0,083 | -0,85 | -2,74 |
| LTG | LowQC | 2 |  | 2,02 | 0,054 | 2,67 | 0,75 |  |  | 1,99 | 0,013 | 0,68 | -0,72 |
|  | MediumQC | 10 |  | 10,14 | 0,007 | 0,07 | 1,44 |  |  | 10,65 | 0,303 | 2,85 | 6,51 |
|  | HighQC | 20 |  | 20,05 | 0,337 | 1,68 | 0,26 |  |  | 18,84 | -0,496 | -2,63 | -5,79 |
| LEV | LowQC | 5 |  | 5,02 | 0,050 | 1,00 | 0,41 |  |  | 5,01 | 0,148 | 2,95 | 0,11 |
|  | MediumQC | 50 |  | 47,39 | 0,943 | 1,99 | -5,23 |  |  | 48,71 | -0,217 | -0,45 | -2,59 |
|  | HighQC | 80 |  | 75,06 | -1,054 | -1,40 | -6,17 |  |  | 81,99 | 2,459 | 3,00 | 2,48 |

|  |  |  |  | Autosampler Stability 24h at 4 °C (n=3) |  |  |  |  | Autosampler Stability 5 days at 4 °C (n=3) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Nominal conc. (μg/mL) |  | Average | Std. Dev. | %CV | %RE |  | Average | Std. Dev. | %CV | %RE |
| CBZ | LowQC | 2 |  | 2,01 | 0,092 | 4,58 | 0,25 |  | 1,82 | -0,154 | -8,47 | -9,00 |
|  | MediumQC | 10 |  | 9,65 | -0,144 | -1,49 | -3,46 |  | 9,67 | -0,881 | -9,11 | -3,31 |
|  | HighQC | 15 |  | 15,48 | 0,271 | 1,75 | 3,20 |  | 17,15 | 1,107 | 6,46 | 14,32 |
| LCM | LowQC | 1 |  | 1,003 | 0,132 | 13,19 | 0,25 |  | 1,04 | 0,062 | 5,99 | 3,90 |
|  | MediumQC | 5 |  | 4,88 | -0,467 | -9,58 | -2,45 |  | 4,70 | -0,612 | -13,02 | -6,00 |
|  | HighQC | 10 |  | 10,22 | 0,296 | 2,89 | 2,18 |  | 10,64 | -1,570 | -14,76 | 6,35 |
| LTG | LowQC | 2 |  | 2,00 | 0,091 | 4,52 | 0,10 |  | 1,93 | 0,267 | 13,86 | -3,60 |
|  | MediumQC | 10 |  | 9,91 | -0,964 | -9,73 | -0,88 |  | 9,41 | 0,059 | 0,63 | -5,91 |
|  | HighQC | 20 |  | 20,16 | 0,419 | 2,08 | 0,78 |  | 20,50 | 1,413 | 6,89 | 2,50 |
| LEV | LowQC | 5 |  | 4,99 | 0,037 | 0,74 | -0,12 |  | 5,10 | 0,110 | 2,15 | 2,01 |
|  | MediumQC | 50 |  | 51,54 | 1,947 | 3,78 | 3,09 |  | 49,18 | 2,606 | 5,30 | -1,64 |
|  | HighQC | 80 |  | 77,63 | -3,227 | -4,16 | -2,96 |  | 83,22 | 0,711 | 0,85 | 4,03 |

|  |  |  |  | 15 days extracted stability at -20 °C (n=3) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Nominal conc. (μg/mL) |  | Average | Std. Dev. | %CV | %RE |
| CBZ | LowQC | 2 |  | 2,02 | 0,042 | 2,06 | 1,17 |
|  | MediumQC | 10 |  | 8,58 | -0,016 | -0,18 | -14,17 |
|  | HighQC | 15 |  | 16,10 | 0,368 | 2,29 | 7,34 |
| LCM | LowQC | 1 |  | 1,027 | 0,023 | 2,27 | 2,65 |
|  | MediumQC | 5 |  | 4,32 | -0,021 | -0,49 | -13,59 |
|  | HighQC | 10 |  | 8,67 | -0,544 | -6,28 | -13,30 |
| LTG | LowQC | 2 |  | 2,01 | 0,006 | 0,28 | 0,55 |
|  | MediumQC | 10 |  | 9,50 | 0,180 | 1,89 | -5,01 |
|  | HighQC | 20 |  | 20,89 | 0,205 | 0,98 | 4,45 |
| LEV | LowQC | 5 |  | 5,02 | 0,052 | 1,03 | 0,43 |
|  | MediumQC | 50 |  | 44,70 | -0,725 | -1,62 | -10,61 |
|  | HighQC | 80 |  | 78,15 | -0,066 | -0,08 | -2,31 |

Figure S1: Selectivity

Figure S2: Carry-over