Supplementary material to:

Analysis of Δ9-tetrahydrocannabinol (Δ9-THC) and cannabidiol (CBD) on hair after single and repeated short in vivo passive exposures to low- and high-Δ9-THC-cannabis

Content of the Supplementary:

- Urine analyses and sample preparation for Δ9-THC-COOH analysis and CBD-COOH analysis

- Gas chromatography-mass spectrometry (GC-MS) method for urine analyses and results of validation

A. Urine analyses and sample preparation for Δ9-THC-COOH analysis and CBD-COOH analysis

To test for Δ9-THC-COOH, 300 μl of potassium hydroxide and 10 μl of the 0.1 mg/ml IS (Δ9-THC-COOH-D3) were added. To a volume of 3 ml of urine. The sample was placed in an oven at 60°C for 15 min.

The sample was then adjusted to pH between 4.5 and 6.5 by adding glacial acetic acid and a mixture of 0.1 M sodium acetate buffer/methanol (95:5), and subjected to solid phase extraction (SPE).

The SPE cartridge was previously conditioned with 1 ml of methanol and 1 ml of 0.1 M sodium acetate buffer/methanol (95:5). After the sample slowly passed through the SPE cartridge, it was washed sequentially with 2 ml of ultrapure water, 2 ml of methanol/water mixture (50:50, v/v). After vacuum drying, the active ingredients of interest were eluted with 2 ml of a mixture of hexane/ethyl acetate (75:25, v/v) with 1% glacial acetic acid.

The eluates were collected and evaporated to dryness under a gentle flow of nitrogen at 45 °C. For the derivatization procedure, 50 μl of BSTFA 1%TMCS was added to the dried residue and shaken for 10 s. The tubes were heated in an oven at 70 °C for 30 min.

A 1 μl aliquot of the resulting solution was injected into the GC-MS system with a previously validated method.

The same extraction process was also used to test for CBD-COOH, but without the initial hydrolysis.

B. Gas chromatography-mass spectrometry (GC-MS) method for urine analyses and results of validation

Samples were analyzed using a Shimadzu gas chromatograph, combined with a QP 2010 Ultra quadrupole mass spectrometer. Data were acquired and analyzed using Shimadzu GCMS Solution. Separation of analytes was achieved using a capillary column Rxi-5Sil MS (30 m × 0.25 mm i.d., 0.25 μm film thickness), supplied by Restek S.r.l.. The carrier gas (helium) was set at total flow rate of 5.4 ml/min. The injection volume was 1 μl in split mode, the injection port, and interface temperatures were set at 290 °C and 270 °C, respectively. The oven temperature was started at 200 °C for 1 min, followed by temperature ramps of 30 °C/min to 250 °C for 0.5 min, 30 °C/min to 280 °C for 6 min, and 30 °C/min to 290 °C held for 9 min. The total separation run time was 19.5 min. The ion source was held at 230 °C. The mass spectrometer was operated with a filament potential of 1.24 KV and an electron energy of 70 eV in the positive electron ionization mode. The selected ion monitoring mode was used with a event time of 0.17 s. Three ions for each analyte and three ions per SI were chosen based on selectivity and abundance to maximize the signal-to-noise ratio in the matrix extracts.

For Δ9-THC-COOH, the method showed linearity (R2> 0.997) in the range 5-250 ng/ml, lower limit of quantification (LLOQ) at the lowest calibration point, limit of detection (LOD) at 1.666 ng/ml and acceptable accuracy (bias% ≤20%) and precision (coefficient of variation% ≤20%).

The method also showed linearity (R2> 0.997) in the range 5-250 ng/ml, LLOQ at the lowest calibration point, LOD at 1.666 ng/ml and acceptable accuracy (bias% ≤20%) and precision (coefficient of variation% ≤20%) for CBD-COOH.