## data Apple
Study | Participants | detail of condition | Age | Age_SD | Age_group | BMI | Female % | Fitzpatrick scale (median) | N_study | Device | Wrist | Outcome | Criterion_measure | Lab_field | Activities | Intensities | relative_absolute | Duration_minutes | Mean_Criterion | SD_Criterion | Mean_Device | SD_Device | Correlation | BA_bias | BA_SD | LOA_L | LOA_U | Sensitivity | Sens_SD | Specificity | Spec_SD | Repeated_measures | N_analysis | Q1_samplesize | Q2_peerreview | Q3_firmware | Q4_multiple devices | Total | Test-Ref or Ref-Test
Helmer 2022 | clinical | Postoperative | 68 | 11.8519 | old | 26.4 | 37.5 | 2 | 112 | Watch 7 | Random | HR | ECG | Lab | Rest | Rest | relative | [5 empty] | 0.98 | 0.36 |  | 0.09 | 0.63 |  |  |  |  | yes | 327 | 0 | 1 | 1 | 0 | 2 | Test-Ref
Khushhal 2024 | clinical | NCD | 51 | 7 | middle |  | 0 | 2 | 41 | Watch 8 | left | HR | Polar | Lab | Cycling | MV | relative | [5 empty] | 0.99 | -0.01 |  | -0.33 | 0.32 |  |  |  |  | no | 41 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Helmer 2024 | clinical | Postoperative | 63 | 13.5 | middle |  | 70 | 2 | 10 | Watch 7 | Random | HR | ECG | Lab | Mix | Light | relative | [6 empty] | -0.73 |  | -5.9 | 4.45 |  |  |  |  | yes | 8057 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Watch 6 | dominant | HR | Polar | Lab | Rest | Rest | relative | [6 empty] | 0.08 |  | -1.1 | 1.25 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Watch 6 | dominant | HR | Polar | Lab | Walking Trdm | MV | relative | [6 empty] | 0.07 |  | -7.48 | 7.62 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Watch 6 | dominant | HR | Polar | Lab | Running Trdm | MV | relative | [6 empty] | -1.22 |  | -14.02 | 11.57 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Watch 6 | dominant | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -1.3 |  | -6.87 | 4.28 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 59 | Watch 6 | dominant | HR | Polar | Lab | cycling | MV | relative | [6 empty] | -3.15 |  | -12.44 | 6.14 |  |  |  |  | no | 59 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Le 2022 | healthy |  | 23 | 2.2 | young | 22.1 | 50 |  | 20 | Watch 6 | random | EE | IC | Field | walking | Light | relative | [6 empty] | 20.4 |  | -1.6 | 42.5 |  |  |  |  | no | 20 | 0 | 1 | 0 | 0 | 1 | Test-Ref
Le 2022 | healthy |  | 23 | 2.2 | young | 22.1 | 50 |  | 20 | Watch 6 | random | EE | IC | Field | running | MV | relative | [6 empty] | 25.6 |  | -5.7 | 57 |  |  |  |  | no | 20 | 0 | 1 | 0 | 0 | 1 | Test-Ref
Sun 2023 | healthy |  | 22.5 | 1.8 | young | 22.9 | 0 |  | 11 | Watch 6 | random | EE | IC | Lab | Running Trdm | MV | relative | [6 empty] | -80.77 |  | -157.4 | -4.17 |  |  |  |  | no | 11 | 0 | 1 | 0 | 0 | 1 | Ref-test
Sun 2023 | healthy |  | 22.5 | 1.8 | young | 22.9 | 0 |  | 11 | Watch 6 | random | EE | IC | Lab | running | MV | relative | [6 empty] | -67.57 |  | -135 | -0.14 |  |  |  |  | no | 11 | 0 | 1 | 0 | 0 | 1 | Ref-test
Kim 2023 | clinical | Coronary artery disease | 60.9 | 7.9 | middle | 25.5565 | 18.1818 |  | 44 | Watch 7 | random | HR | ECG | Lab | Running Trdm | MV | relative | [6 empty] | -0.11 |  | -3.33 | 3.22 |  |  |  |  | yes | 1054 | 0 | 1 | 0 | 1 | 3 | Ref-test
O'Grady 2024 | healthy |  | 24.6 | 8.2 | young | 23.65 | 56.41 |  | 39 | Watch 9 | random | HRV | Polar | Field | Rest | Rest | [7 empty] | -8.31 |  | -11 | 5.6 |  |  |  |  | yes | 316 | 0 | 1 | 0 | 1 | 4 | Test-Ref
O'Grady 2024 | healthy |  | 24.6 | 8.2 | young | 23.65 | 56.41 |  | 39 | Watch 9 | random | HR | Polar | Field | Rest | Rest | [7 empty] | -0.08 |  | -0.78 | 0.93 |  |  |  |  | yes | 316 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -5.7 |  | -28.9 | 17.5 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -5.8 |  | -29.9 | 18.3 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | recovery | MV | relative | [6 empty] | 1.7 |  | -4.9 | 8.2 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -7.1 |  | -20.5 | 5.8 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -8.9 |  | -32.4 | 14.7 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | recovery | MV | relative | [6 empty] | 1.8 |  | -15.8 | 19.4 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -13.9 |  | -47.1 | 19.3 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -20.2 |  | -70.3 | 29.8 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | recovery | MV | relative | [6 empty] | -0.4 |  | -20.3 | 19.5 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -6.1 |  | -24.5 | 12.4 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -4.1 |  | -16.7 | 8.5 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | recovery | MV | relative | [6 empty] | 1.3 |  | -4.9 | 7.5 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -16.2 |  | -53.8 | 21.4 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -22 |  | -61.7 | 17.7 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Støve 2023 | healthy |  | 24.5 | 4 | young | 21.1 | 55 |  | 29 | Watch 6 | random | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -0.5 |  | -18.3 | 17.3 |  |  |  |  | no | 29 | 0 | 1 | 0 | 1 | 4 | Test-Ref
Ho 2022 | healthy |  | 29.3 | 3.6 | young | 21.2 | 50 |  | 30 | Watch 6 | left | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | 1.27 |  | -1.35 | 3.89 |  |  |  |  | no | 30 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Ho 2022 | healthy |  | 29.3 | 3.6 | young | 21.2 | 50 |  | 30 | Watch 6 | left | HR | ECG | Lab | Cycling | MV | relative | [6 empty] | 0.69 |  | -1.87 | 3.26 |  |  |  |  | no | 30 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Ho 2022 | healthy |  | 29.3 | 3.6 | young | 21.2 | 50 |  | 30 | Watch 6 | left | HR | ECG | Lab | Cycling | MV | relative | [6 empty] | 0.4 |  | -2.75 | 3.56 |  |  |  |  | no | 30 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Ho 2022 | healthy |  | 29.3 | 3.6 | young | 21.2 | 50 |  | 30 | Watch 6 | left | HR | ECG | Lab | Cycling | MV | relative | [6 empty] | -0.01 |  | -4.26 | 4.43 |  |  |  |  | no | 30 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Giggins 2021 | healthy |  | 30.4 | 8 | young | 24.67 | 50 | 3 | 10 | Watch 5 | left | HR | Polar | Field | Daily activities | Light | relative | [6 empty] | -5.8 |  | -32.5 | 20.9 |  |  |  |  | yes | 57 | 0 | 1 | 1 | 1 | 3 | test-ref
Giggins 2021 | healthy |  | 30.4 | 8 | young | 24.67 | 50 | 3 | 10 | Watch 5 | left | HR | Polar | Field | Running Trdm | MV | relative | [6 empty] | -5.5 |  | -19.6 | 8.6 |  |  |  |  | yes | 406 | 0 | 1 | 1 | 1 | 3 | test-ref
Giggins 2021 | healthy |  | 30.4 | 8 | young | 24.67 | 50 | 3 | 10 | Watch 5 | left | HR | Polar | Field | Rest | Rest | relative | [6 empty] | -7.5 |  | -22.3 | 15.3 |  |  |  |  | yes | 101 | 0 | 1 | 1 | 1 | 3 | test-ref
Viciana 2022 | healthy |  | 14.7 | 1.7 | young | 21.2 | 48.21 |  | 56 | Watch 5 |  | Step | Video | Lab | Walking | Light | relative | [6 empty] | 1.9 |  | -4.8 | 8.6 |  |  |  |  | no | 56 | 1 | 1 | 0 | 0 | 2 | ref-test
Viciana 2022 | healthy |  | 14.7 | 1.7 | young | 21.2 | 48.21 |  | 56 | Watch 5 |  | Step | Video | Lab | Walking | Light | relative | [6 empty] | 3.1 |  | -6.9 | 13.1 |  |  |  |  | no | 56 | 1 | 1 | 0 | 0 | 2 | ref-test
Viciana 2022 | healthy |  | 14.7 | 1.7 | young | 21.2 | 48.21 |  | 56 | Watch 5 |  | Step | Video | Lab | Walking | MV | relative | [6 empty] | 6.7 |  | -14.5 | 27.9 |  |  |  |  | no | 56 | 1 | 1 | 0 | 0 | 2 | ref-test
Viciana 2022 | healthy |  | 14.7 | 1.7 | young | 21.2 | 48.21 |  | 56 | Watch 5 |  | Step | Video | Lab | Walking | MV | relative | [6 empty] | 0.2 |  | -4.1 | 4.5 |  |  |  |  | no | 56 | 1 | 1 | 0 | 0 | 2 | ref-test
Robbins 2024 | healthy |  | 32.9 |  | young |  | 57.1 |  | 29 | Watch 8 | Random | Sleep (TST) | PSG | Lab |  |  | relative | [6 empty] | 8 |  | -33 | 48 |  |  |  |  | no | 29 | 0 | 1 | 0 | 0 | 1 | Test-Ref
Robbins 2024 | healthy |  | 32.9 |  | young |  | 57.1 |  | 33 | Watch 8 | Random | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 97 | 2 | 52.4 | 21.5 | no | 33 | 0 | 1 | 0 | 0 | 1 | Test-Ref
Miller2022 | healthy |  | 25.4 | 5.9 | young |  | 49.6 |  | 53 | Watch 6 | random | Sleep (TST) | PSG | Lab |  |  | relative | [6 empty] | 39.5 | 41.5 | [6 empty] | no | 53 | 1 | 1 | 0 | 0 | 2 | Test-Ref
Miller2022 | healthy |  | 25.4 | 5.9 | young |  | 49.6 |  | 53 | Watch 6 | random | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 97 |  | 26 |  | no | 53 | 1 | 1 | 0 | 0 | 2
Miller2022 | healthy |  | 25.4 | 5.9 | young |  | 49.6 |  | 53 | Watch 6 | random | HR | ECG | Lab | Sleeping |  | relative | [6 empty] | 0.5 | 2.1 | -3.5 | 4.5 |  |  |  |  | no | 53 | 1 | 1 | 0 | 0 | 2 | Test-Ref
Lee2023 | clinical | sleep disorders | 43.59 | 14.1 | middle | 23.9 | 48 |  | 26 | Watch 8 | Random | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 96 |  | 45 |  | no | 26 | 0 | 1 | 0 | 0 | 1