## data_r_4
Study | Participants | detail of condition | Age | Age_SD | Age_group | BMI | Female % | Fitzpatrick scale (median) | N_study | Device | Wrist | Outcome | Criterion_measure | Lab_field | Activities | Intensities | relative_absolute | Duration_minutes | Mean_Criterion | SD_Criterion | Mean_Device | SD_Device | Correlation | BA_bias | BA_SD | LOA_L | LOA_U | Sensitivity | Sens_SD | Specificity | Spec_SD | Repeated_measures | N_analysis | Q1_samplesize | Q2_peerreview | Q3_firmware | Q4_multiple devices | Total | Test-Ref or Ref-Test
Helmer 2022 | clinical | postoperative patients after moderate to major surgery | 68 | 11.8519 | old | 26.4 | 37.5 | 2 | 112 | Sense | Random | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | 0.77 |  | 0.28 | 1.26 |  |  |  |  | yes | 308 | 0 | 1 | 1 | 0 | 2 | Test-Ref
Meza 2023 | clinical | stroke | 79.4 | 10.1 | old | 26.3 |  |  | 70 | Charge 5 | non-dominant | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | -10.179 |  | -53.216 | 32.858 |  |  |  |  | yes | 526 | 1 | 1 | 0 | 0 | 2 | Ref-Test
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Rest | Rest | relative | [6 empty] | -2 |  | -12 | 9 |  |  |  |  | yes | 90 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Walking | MV | relative | [6 empty] | -4 |  | -34 | 27 |  |  |  |  | yes | 90 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Recovery | Recovery | relative | [6 empty] | 2 |  | -13 | 17 |  |  |  |  | yes | 90 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Rest | Rest | relative | [6 empty] | 0 |  | -8 | 9 |  |  |  |  | yes | 45 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Alternate Stepping | MV | relative | [6 empty] | -10 |  | -39 | 18 |  |  |  |  | yes | 45 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Recovery | Recovery | relative | [6 empty] | 0 |  | -11 | 12 |  |  |  |  | yes | 45 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Rest | Rest | relative | [6 empty] | 1 |  | -12 | 13 |  |  |  |  | yes | 48 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Chair Rise | MV | relative | [6 empty] | -13 |  | -26 | 0 |  |  |  |  | yes | 48 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | HR | ECG | Field | Recovery | Recovery | relative | [6 empty] | 0 |  | -6 | 6 |  |  |  |  | yes | 48 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | Distance | Meter wheel | Field | Walking | MV | relative | [6 empty] | -77 |  | -205 | 52 |  |  |  |  | yes | 90 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | Distance | Meter wheel | Field | Walking | MV | relative | [6 empty] | -55 |  | -145 | 36 |  |  |  |  | yes | 45 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | Steps | Hand tally counter | Field | Walking | MV | relative | [6 empty] | -142 |  | -275 | -8 |  |  |  |  | yes | 90 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jamieson 2024 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 | 2 | 16 | Sense | Random | Steps | Hand tally counter | Field | Walking | MV | relative | [6 empty] | -75 |  | -216 | 66 |  |  |  |  | yes | 90 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Cuegniez 2025 | healthy |  | 21.1 | 5.8 | young | 21.9 | 44 | 2.9 | 7 | Charge 4 | non-dominant | HR | Polar | Field | Badminton | MV | relative | [6 empty] | -16.5 |  | -68.2 | 35.2 |  |  |  |  | yes | 904 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Cuegniez 2025 | healthy |  | 21.1 | 5.8 | young | 21.9 | 44 | 2.9 | 2 | Charge 4 | non-dominant | HR | Polar | Field | Tennis | MV | relative | [6 empty] | -6.2 |  | -37.2 | 24.8 |  |  |  |  | yes | 294 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Cuegniez 2025 | healthy |  | 21.1 | 5.8 | young | 21.9 | 44 | 2.9 | 5 | Charge 4 | non-dominant | HR | Polar | Field | Orienteering running | MV | relative | [6 empty] | -8.6 |  | -43.3 | 26 |  |  |  |  | yes | 421 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Cuegniez 2025 | healthy |  | 21.1 | 5.8 | young | 21.9 | 44 | 2.9 | 3 | Charge 4 | non-dominant | HR | Polar | Field | running | MV | relative | [6 empty] | 0.1 |  | -10.7 | 10.9 |  |  |  |  | yes | 845 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Cuegniez 2025 | healthy |  | 21.1 | 5.8 | young | 21.9 | 44 | 2.9 | 2 | Charge 4 | non-dominant | HR | Polar | Field | cycling | MV | relative | [6 empty] | 4.8 |  | -27.3 | 36.8 |  |  |  |  | yes | 1103 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Cuegniez 2025 | healthy |  | 21.1 | 5.8 | young | 21.9 | 44 | 2.9 | 12 | Charge 4 | non-dominant | HR | Polar | Field | soccer | MV | relative | [6 empty] | -16.5 |  | -59.6 | 26.7 |  |  |  |  | yes | 1080 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Jachymek 2021 | healthy |  | 28 |  | young | 23.5 | 32.26 |  | 31 | Charge 4 | left | HR | ECG | Lab | Trdm | MV | relative | [6 empty] | -9.348 |  | -46.11 | 27.41 |  |  |  |  | yes | 439 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Sense | dominant | HR | Polar | Lab | Rest | Rest | relative | [6 empty] | 1.14 |  | -13.43 | 15.7 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Sense | dominant | HR | Polar | Lab | Trdm | MV | relative | [6 empty] | -1.26 |  | -13.44 | 10.91 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Sense | dominant | HR | Polar | Lab | Trdm | MV | relative | [6 empty] | -3.76 |  | -16.73 | 9.22 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 60 | Sense | dominant | HR | Polar | Lab | resistance excercises | MV | relative | [6 empty] | -6.51 |  | -19.48 | 6.46 |  |  |  |  | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hajj-Boutros 2023 | healthy |  | 24.9 | 3 | young | 23.1 | 50 |  | 59 | Sense | dominant | HR | Polar | Lab | cycling | MV | relative | [6 empty] | -7.86 |  | -19.89 | 4.17 |  |  |  |  | no | 59 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Hermans 2024 | healthy |  | 68 | 7 | old | 28 | 20 |  | 20 | Charge 4 | non-dominant | HR | Polar | Field | Rest | Rest | relative | [6 empty] | 3 |  | -10 | 15 |  |  |  |  | no | 20 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Hermans 2024 | clinical | COPD | 70 | 7 | old | 27 | 6.6667 |  | 25 | Charge 4 | non-dominant | HR | Polar | Field | Rest | Rest | relative | [6 empty] | -0.3 |  | -9 | 8 |  |  |  |  | no | 25 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Hermans 2024 | healthy |  | 68 | 7 | old | 28 | 20 |  | 20 | Charge 4 | non-dominant | HRV | Polar | Field | Rest | Rest | relative | [6 empty] | -7 |  | -30 | 16 |  |  |  |  | no | 20 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Hermans 2024 | clinical | COPD | 70 | 7 | old | 27 | 6.6667 |  | 25 | Charge 4 | non-dominant | HRV | Polar | Field | Rest | Rest | relative | [6 empty] | -3 |  | -22 | 15 |  |  |  |  | no | 25 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Simon-Vicente 2024 | clinical | Huntington's disease | 55.7 | 11.4 | middle | 24.8 | 50 |  | 14 | Charge 4 | dominant | EE | IC | Lab | Walking Trdm | MV | relative | [6 empty] | 3.169 |  | 0.033 | 6.304 |  |  |  |  | no | 14 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Simon-Vicente 2024 | clinical | Huntington's disease | 55.7 | 11.4 | middle | 24.8 | 50 |  | 14 | Charge 4 | dominant | EE | IC | Lab | Walking Trdm | MV | relative | [6 empty] | 2.9 |  | -0.321 | 6.121 |  |  |  |  | no | 14 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Simon-Vicente 2024 | clinical | Huntington's disease | 55.7 | 11.4 | middle | 24.8 | 50 |  | 14 | Charge 4 | dominant | EE | IC | Lab | Walking Trdm | MV | relative | [6 empty] | 2.59 |  | -0.962 | 6.143 |  |  |  |  | yes | 144 | 0 | 1 | 0 | 1 | 2 | Ref-Test
Stutz 2024 | healthy |  | 64 |  | middle | 24.4 | 48 |  | 60 | Charge 5 | dominant | EE | Oxycon | Lab | Mix | Light | relative | [6 empty] | -0.64 |  | -4.17 | 2.9 |  |  |  |  | no | 60 | 1 | 1 | 0 | 0 | 2 | Test-Ref
Lim2023 | healthy |  | 39 | 8 | young | 23.6 | 55.6 |  | 9 | Inspire 2 | Random | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 93.9 |  | 13.1 |  | yes | 18 | 0 | 1 | 0 | 1 | 2
Dong2022 | clinical | insomnia | 48.8 | 2.1 | middle |  | 54.05 |  | 37 | Charge 4 | non-dominant | Sleep (TST) | PSG | Lab |  |  | relative | [6 empty] | -11 |  | -99.3 | 77.2 |  |  |  |  | no | 37 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Dong2022 | clinical | insomnia | 48.8 | 2.1 | middle |  | 54.05 |  | 37 | Charge 4 | non-dominant | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 89.9 | 4 | 62.2 | 26.2 | no | 37 | 0 | 1 | 0 | 1 | 2
Ong2023 | healthy |  | 38.47 | 15.05 | young |  | 56.06 |  | 60 | Sense | non-dominant | Sleep (TST) | PSG | Lab |  |  | relative | [6 empty] | 5.97 | 40.66 | [6 empty] | no | 60 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Ong2023 | healthy |  | 38.47 | 15.05 | young |  | 56.06 |  | 60 | Sense | non-dominant | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 94.46 | 3.87 | 67.77 | 20.76 | no | 60 | 1 | 1 | 1 | 0 | 3
Willoughby2024 | healthy |  | 39.56 | 15.51 | young | 23.84 | 57.14 |  | 60 | Sense | non-dominant | Sleep (TST) | PSG | Lab |  |  | relative | [6 empty] | -12.65 | 33.69 | [6 empty] | no | 50 | 1 | 1 | 1 | 0 | 3 | Test-Ref
Willoughby2024 | healthy |  | 39.56 | 15.51 | young | 23.84 | 57.14 |  | 50 | Sense | non-dominant | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 92.32 | 5.17 | 78.68 | 17.71 | no | 50 | 1 | 1 | 1 | 0 | 3
Grandner2023 | healthy |  | 23.4 | 3.8 | young | 23.4 | 44.4 |  | 36 | Charge 4 | left | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 89 | 9 | 49 | 18 | no | 36 | 0 | 1 | 1 | 0 | 2
Robbins 2024 | healthy |  | 32.9 |  | young |  | 57.1 |  | 33 | Sense 2 | Random | Sleep (TST) | PSG | Lab |  |  | relative | [6 empty] | -3 |  | -63 | 57 |  |  |  |  | no | 33 | 0 | 1 | 0 | 0 | 1 | Test-Ref
Robbins 2024 | healthy |  | 32.9 |  | young |  | 57.1 |  | 33 | Sense 2 | Random | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 95 | 3 | 67.7 | 23.1 | no | 33 | 0 | 1 | 0 | 0 | 1 | Test-Ref
Li 2024 | clinical | postoperative | 52.2 | 13.6 | middle | 22.5 | 40 |  | 45 | Fitbit Inspire HR |  | Steps | Hand tally counter | Lab | Walking | MV | relative | [6 empty] | -205.1 |  | -489.7 | 79.5 |  |  |  |  | no | 45 | 0 | 1 | 0 | 0 | 1 | tesr-ref
Colonna 2023 | clinical | Parkinson | 65.32 | 5.76 | old |  | 45.45 |  | 11 | Charge 4 | side less affected by PD | HR | Polar | Lab | Rest | Rest | relative | [6 empty] | 1.68 |  | -21.52 | 18.8 |  |  |  |  | yes | 596 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Colonna 2023 | clinical | Parkinson | 65.32 | 5.76 | old |  | 45.45 |  | 11 | Charge 4 | side less affected by PD | HR | Polar | Lab | Maximal exercise | MV | relative | [6 empty] | 6.29 |  | -22.02 | 36.2 |  |  |  |  | yes | 596 | 0 | 1 | 0 | 1 | 2 | Test-Ref
Ogasawara2023 | clinical | psychiatric disorders | 48.1 | 17.2 | middle |  | 61.5 |  | 52 | Sense | non-dominant | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 95 | 7 | 45 | 23 | no | 52 | 1 | 1 | 0 | 0 | 2
Lee2023 | clinical | sleep disorders | 43.59 | 14.1 | middle | 23.9 | 48 |  | 26 | Sense 2 | Random | Sleep (epoch) | PSG | Lab |  |  | relative | [10 empty] | 96 |  | 27 |  | no | 26 | 0 | 1 | 0 | 0 | 1

## discarded
Reason | Study | Participants | detail of condition | Age | Age_SD | Age_group | BMI | Female % | Fitzpatrick scale (median) | N_study | Device | Wrist | Outcome | Criterion_measure | Lab_field | Activities | Intensities | relative_absolute | Duration_minutes | Mean_Criterion | SD_Criterion | Mean_Device | SD_Device | Correlation | BA_bias | BA_SD | LOA_L | LOA_U | Repeated_measures | N_analysis | Q1_samplesize | Q2_peerreview | Q3_firmware | Q4_multiple devices | Total | Test-Ref or Ref-Test
| Alyson 2025 | clinical | Fontan physiology | 17.7 | 3.4815 | young |  | 44.7 |  | 47 | Charge 4 | right | HR | ECG | Lab | Rest | Rest | relative | [10 empty] | no | 47 | 0 | 1 | 0 | 1 | 2 | Ref-Test
| Alyson 2025 | clinical | Fontan physiology | 17.7 | 3.4815 | young |  | 44.7 |  | 47 | Charge 4 | right | HR | ECG | Lab | Maximal exercise | MV | relative | [10 empty] | no | 47 | 0 | 1 | 0 | 1 | 2 | Ref-Test
| Alyson 2025 | clinical | Fontan physiology | 17.7 | 3.4815 | young |  | 44.7 |  | 47 | Charge 4 | right | HR | ECG | Lab | Recovery | Rest | relative | [10 empty] | no | 47 | 0 | 1 | 0 | 1 | 2 | Ref-Test
| Nissen 2022 | healthy |  | 24.2 | 4.6 | young |  | 43 |  | 23 | Fitbit Charge 4 and Samsung Active 2 | random | HR | ECG | Lab | Mix | Light | relative | [6 empty] | -1.66 |  | -26.75 | 23.43 | yes | da escludere? | 0 | 1 | 0 | 1 | 2 |  | Very difficult to understand the N_Analysis. There are several measurements but no values or method to derive them are present
only Bland-Altman plots, no value | Jamieson 2022 | healthy |  | 27 | 2.2222 | young | 23.9 | 56 |  | 16 | Sense 2
N_analysis?? | Gagnon 2022 | healthy |  | 26.8 | 8.5 | young |  | 58 |  | 34 | fitbit versa 2 | nondominant | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | -0.27 |  | -8.7 | 8.16 |  |  | 0 | 1 | 1 | 1 | 3
N_analysis?? | Gagnon 2022 | healthy |  | 26.8 | 8.5 | young |  | 58 |  | 34 | fitbit versa 2 | nondominant | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | -6.7 |  | -20.36 | 6.97 |  |  | 0 | 1 | 1 | 1 | 3
N_analysis?? | Gagnon 2022 | healthy |  | 26.8 | 8.5 | young |  | 58 |  | 34 | fitbit versa 2 | nondominant | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | 0.65 |  | -11.08 | 12.38 |  |  | 0 | 1 | 1 | 1 | 3
N_analysis?? | Gagnon 2022 | healthy |  | 26.8 | 8.5 | young |  | 58 |  | 34 | fitbit versa 2 | nondominant | HR | ECG | Lab | Rest | Rest | relative | [6 empty] | -1.32 |  | -12.95 | 10.31 |  |  | 0 | 1 | 1 | 1 | 3
CANNOT compute BIAS_SD | Lim2023 | healthy |  | 39 | 8 | young | 23.6 | 55.6 |  | 9 | Inspire 2 | Random | Sleep (TST) | PSG | Lab |  |  | relative |  | 346.42 | 73.89 | 364.33 | 79.98 |  | 17.91 |  |  |  | yes | 18 | 0 | 1 | 0 | 1 | 2 | Test-Ref
CANNOT compute BIAS_SD | Ogasawara2023 | clinical | psychiatric disorders | 48.1 | 17.2 | middle |  | 61.5 |  | 52 | Sense | non-dominant | Sleep (TST) | PSG | Lab |  |  | relative |  | 456.19 | 83.16 | 491.9 | 40.47 | [6 empty] | 52 | 1 | 1 | 0 | 0 | 2 | Test-Ref
Note on Q11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: anticipation
Note on Q12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Oral
Note on Q13: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Arithmetic