Suppl. Table 1. 

Evaluation of cannabis STR typing of three replicates of sample 1 by the laboratories participating in the collaborative exercise.

| Lab | Number of typed loci | Number of loci with concordant results | Missing one allele at heterozygous locus | Number of loci with wrong results | Number of loci with no results | Percentage of success |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 2 | 39 | 33 | 6 | 2 | 0 | 84,6% |
| 3 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 4 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 5 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 6 | 39 | 29 | 1 | 1 | 9 | 74,4% |
| 7 | 39 | 34 | 5 | 5 | 0 | 87,2% |
| 8 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 9 | 39 | 36 | 3 | 3 | 0 | 92,3% |
| 10 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 11 | 39 | 37 | 2 | 2 | 0 | 94,9% |
| 12 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 13 | 39 | 29 | 4 | 4 | 6 | 74,4% |
| 14 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 15 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 16 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 17 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 18 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 19 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 20 | 39 | 39 | 0 | 0 | 0 | 100,0% |

Mean 95.4%

Suppl. Table 2.
 
Evaluation of cannabis STR typing of three replicates of sample 2 by the laboratories participating in the collaborative exercise.

| Lab | Number of typed loci | Number of loci with concordant results | Missing one allele at heterozygous locus | Number of loci with wrong results | Number of loci with no results | Percentage of success |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 2 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 3 | 39 | 36 | 3 | 3 | 0 | 92,3% |
| 4 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 5 | 39 | 32 | 7 | 7 | 7 | 82,1% |
| 6 | 39 | 27 | 12 | 12 | 13 | 69,2% |
| 7 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 8 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 9 | 13 | 13 | 0 | 0 | 0 | 100,0% |
| 10 | 26 | 26 | 0 | 0 | 0 | 100,0% |
| 11 | 26 | 26 | 0 | 0 | 0 | 100,0% |
| 12 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 13 | 39 | 36 | 3 | 3 | 0 | 92,3% |
| 14 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 15 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 16 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 17 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 18 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 19 | 39 | 39 | 0 | 0 | 0 | 100,0% |
| 20 | 39 | 39 | 0 | 0 | 0 | 100,0% |

Mean 96,8%.

Suppl. Table 3.

Sample 1 - Stutter Analysis.

| Locus | Allele | A L M P H (RFU) | S E M | min (RFU) | max (RFU) | M P H PAS(RFU) | -1 Stutters | -1 Stutter % (RFU) | S E M | +1 Stutters | +1 Stutter % (RFU) | S E M |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ANUCS501 | 6 | 1117 | 1531 | 81 | 6636 | 6636 | 1 | 1,26 |  |  |  |  |
| 9269 | 6 | 587 | 763 | 51 | 3249 |  |  |  |  |  |  |  |
| 4910 | 4 | 1059 | 1534 | 60 | 6583 |  |  |  |  |  |  |  |
| 4910 | 10 | 801 | 1044 | 48 | 4406 | 1448 | 4 | 3,13 | 0,34 |  |  |  |
| 5159 | 4 | 2118 | 1355 | 410 | 4876 | 767 | 10 | 6,58 | 3,16 |  |  |  |
| 5159 | 4.2 | 1413 | 1092 | 191 | 3855 |  |  |  |  |  |  |  |
| ANUCS305 | 8 | 1056 | 1701 | 61 | 7456 | 1667 | 3 | 2,17 | 0,59 |  |  |  |
| 9043 | 5 | 1733 | 1744 | 155 | 7704 |  |  |  |  |  |  |  |
| 9043 | 6 | 1793 | 1925 | 171 | 8765 |  |  |  |  |  |  |  |
| B05 | 8 | 9756 | 7510 | 2086 | 24139 | 2086 | 20 | 3,75 | 1,16 | 14 | 3,04 | 2,6 |
| 1528 | 6 | 2189 | 1764 | 391 | 7347 |  |  |  |  |  |  |  |
| 1528 | 7 | 2245 | 1433 | 305 | 5737 |  |  |  |  |  |  |  |
| 3735 | 3 | 740 | 1146 | 72 | 5099 |  |  |  |  |  |  |  |
| 3735 | 9 | 659 | 915 | 63 | 3852 | 401 | 5 | 8,98 | 5,05 |  |  |  |
| CS1 | 23 | 2478 | 2046 | 302 | 7978 | 1011 | 11 | 5,67 | 5,06 | 3 | 4,36 | 0,66 |
| CS1 | 28 | 1993 | 1242 | 486 | 4303 | 486 | 18 | 8,72 | 5,79 |  |  |  |
| D02 | 7 | 7658 | 5942 | 839 | 21053 | 2456 | 18 | 4,16 | 1,23 | 3 | 3,35 | 2,78 |
| C11 | 14 | 5795 | 4047 | 1111 | 13630 | 1111 | 19 | 6,49 | 3 | 3 | 5,68 | 1,5 |
| H06 | 8 | 2127 | 2135 | 231 | 8257 | 1014 | 7 | 5,77 | 2,89 |  |  |  |

ALMPH (Average of Laboratories Mean Peak Height in RFU): Allele height mean of each laboratory allele height means; SEM: Standard error value; Min: Lower value of RFU where that allele was obtained; Max: Maximum value of RFU where that allele was obtained; MPHPAS (Minimum Peak Height of Parental Allele with a Stutter): Lower value in RFU obtained for a parental allele with stutter, -1 Stutters: Number of -1 stutters found in that allele; -1 Stutter %: Percentage of -1 stutter for that locus, +1 Stutters: Number of +1 stutters found in that allele; +1 Stutter %: Percentage of -1 stutter for that locus.

Suppl. Table 4.

Sample 2 - Stutter Analysis

| Locus | Allele | A L M P H (RFU) | S E M | min (RFU) | max (RFU) | M P HPA S (RFU) | -1 Stutters | -1 Stutter % (RFU) | S E M | +1 Stutters | +1 Stutter % (RFU) | S E M |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ANUCS501 | 5 | 600 | 740 | 44 | 1917 |  |  |  |  |  |  |  |
| ANUCS501 | 6 | 479 | 583 | 44 | 1532 |  |  |  |  |  |  |  |
| 9269 | 5.3 | 219 | 168 | 47 | 447 |  |  |  |  |  |  |  |
| 9269 | 6 | 201 | 164 | 44 | 474 |  |  |  |  |  |  |  |
| 4910 | 4 | 1752 | 1818 | 179 | 5179 |  |  |  |  |  |  |  |
| 5159 | 4 | 3202 | 2250 | 949 | 10824 | 1288 | 4 | 4,3 | 1,28 |  |  |  |
| 5159 | 4.2 | 1973 | 1328 | 405 | 5388 |  |  |  |  |  |  |  |
| ANUCS305 | 6 | 477 | 546 | 41 | 1625 |  |  |  |  |  |  |  |
| ANUCS305 | 8 | 511 | 528 | 48 | 1546 |  |  |  |  |  |  |  |
| 9043 | 5 | 3341 | 3100 | 540 | 9702 | 1890 |  |  |  | 3 | 2,47 | 1,34 |
| B05 | 8 | 6890 | 5178 | 6890 | 6890 | 1343 | 19 | 3,45 | 0,82 |  |  |  |
| B05 | 9 | 6836 | 5044 | 1275 | 22623 | 4149 |  |  |  | 1 | 1,7 |  |
| 1528 | 7 | 5169 | 4837 | 893 | 21510 | 3938 | 4 | 2,48 | 1,73 |  |  |  |
| 3735 | 4 | 1148 | 1335 | 51 | 3549 | 1952 | 4 | 2,4 | 1,1 |  |  |  |
| CS1 | 23 | 3404 | 2569 | 504 | 10311 | 1024 | 13 | 6,22 | 3,73 | 1 | 1,31 |  |
| CS1 | 27 | 3111 | 2942 | 714 | 14049 | 1003 | 14 | 3,2 | 0,73 | 5 | 3,47 | 0,49 |
| D02 | 6 | 3655 | 3100 | 497 | 10928 | 497 | 10 | 1,58 | 0,47 |  |  |  |
| D02 | 8 | 3685 | 2928 | 669 | 10694 | 669 | 16 | 5,65 | 3,04 |  |  |  |
| C11 | 13 | 3144 | 2559 | 614 | 11134 | 1039 | 14 | 4,36 | 0,62 |  |  |  |
| C11 | 14 | 2996 | 2583 | 443 | 11071 |  |  |  |  |  |  |  |
| H06 | 8 | 1311 | 1285 | 96 | 4486 | 697 | 2 | 3,41 | 1,17 |  |  |  |

ALMPH (Average of Laboratories Mean Peak Height in RFU): Allele height mean of each laboratory allele height means; SEM: Standard error value; Min: Lower value of RFU where that allele was obtained; Max: Maximum value of RFU where that allele was obtained; MPHPAS (Minimum Peak Height of Parental Allele with a Stutter): Lower value in RFU obtained for a parental allele with stutter, -1 Stutters: Number of -1 stutters found in that allele; -1 Stutter %: Percentage of -1 stutter for that locus, +1 Stutters: Number of +1 stutters found in that allele; +1 Stutter %: Percentage of -1 stutter for that locus.

Suppl. Table 5.

Positive control - Stutter Analysis

| Locus | Allele | A L M P H (RFU) | S E M | min (RFU) | max (RFU) | M P HPA S (RFU) | -1 Stutters | -1 Stutter % (RFU) | S E M | +1 Stutters | +1 Stutter % (RFU) | S E M |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ANUCS501 | 4 | 602 | 779 | 71 | 3042 |  |  |  |  |  |  |  |
| ANUCS501 | 6 | 503 | 639 | 50 | 2164 | 1230 | 3 | 7,46 | 1,74 |  |  |  |
| 9269 | 6 | 2024 | 2023 | 299 | 7498 |  |  |  |  | 3 | 3,68 | 3 |
| 4910 | 4 | 930 | 921 | 127 | 3671 |  |  |  |  |  |  |  |
| 4910 | 10 | 777 | 727 | 73 | 2502 | 931 | 3 | 4,32 | 0,93 |  |  |  |
| 5159 | 4.2 | 2230 | 1377 | 544 | 5220 |  |  |  |  |  |  |  |
| ANUCS305 | 4 | 480 | 575 | 67 | 2055 |  |  |  |  |  |  |  |
| 9043 | 5 | 1602 | 1196 | 276 | 5059 |  |  |  |  |  |  |  |
| 9043 | 6 | 1540 | 1240 | 340 | 4805 |  |  |  |  |  |  |  |
| B05 | 8 | 5774 | 2979 | 2142 | 12461 | 2142 | 19 | 3,71 | 1,07 |  |  |  |
| B05 | 9 | 5247 | 2857 | 1439 | 10703 | 3502 |  |  |  | 2 | 1,34 | 0,95 |
| 1528 | 7 | 3786 | 2177 | 867 | 9647 | 6440 | 2 | 5,11 | 1,63 |  |  |  |
| 3735 | 7 | 983 | 1063 | 102 | 4317 | 1548 | 4 | 3,37 | 1,91 |  |  |  |
| CS1 | 10 | 5692 | 2984 | 2399 | 12246 | 3353 | 7 | 1,45 | 0,86 |  |  |  |
| CS1 | 29 | 2234 | 1421 | 472 | 5107 | 472 | 15 | 5,32 | 3,9 |  |  |  |
| D02 | 6 | 7302 | 4786 | 57 | 19484 | 4379 | 14 | 1,71 | 1,1 | 5 | 1,41 | 0,72 |
| C11 | 13 | 3020 | 1841 | 894 | 7021 | 1343 | 15 | 4,98 | 1,19 |  |  |  |
| C11 | 14 | 2685 | 1752 | 704 | 7242 | 1391 |  |  |  | 2 | 4,14 | 1,94 |
| H06 | 7 | 887 | 745 | 125 | 3336 |  |  |  |  |  |  |  |
| H06 | 8 | 808 | 613 | 140 | 2377 |  |  |  |  |  |  |  |

ALMPH (Average of Laboratories Mean Peak Height in RFU): Allele height mean of each laboratory allele height means; SEM: Standard error value; Min: Lower value of RFU where that allele was obtained; Max: Maximum value of RFU where that allele was obtained; MPHPAS (Minimum Peak Height of Parental Allele with a Stutter): Lower value in RFU obtained for a parental allele with stutter, -1 Stutters: Number of -1 stutters found in that allele; -1 Stutter %: Percentage of -1 stutter for that locus, +1 Stutters: Number of +1 stutters found in that allele; +1 Stutter %: Percentage of -1 stutter for that locus.

Suppl. Table 6. 
Background data analyzed for each capillary electrophoresis instrument

| Genetic Analyzer | Dye Channel | Maximum Peak Height (RFU) | Average Peak Height (RFU) | Standard Deviation (SD) | Average +3 SD (LOD) | Average +10 SD (LOQ) | Analitical Threshold (RFU) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 310 | Blue (6-FAM) | 42 | 7.0 | 4.3 | 19.8 | 49.7 | 50 |
|  | Green (VIC) | 55 | 7.8 | 3.4 | 18.1 | 42.0 | 42 |
|  | Yellow (NED) | 27 | 10.4 | 3.4 | 20.5 | 44.0 | 44 |
|  | Red (PET) | 38 | 7.9 | 3.0 | 16.9 | 37.8 | 38 |
| 310 | Blue (6-FAM) | 35 | 6.9 | 3.7 | 17.9 | 43.8 | 44 |
|  | Green (VIC) | 35 | 7.7 | 2.7 | 15.9 | 34.8 | 35 |
|  | Yellow (NED) | 35 | 10.5 | 3.5 | 20.9 | 45.2 | 46 |
|  | Red (PET) | 37 | 7.8 | 2.7 | 16.0 | 35.1 | 36 |
| 310 | Blue (6-FAM) | 42 | 7,0 | 4.6 | 20.7 | 52.8 | 53 |
|  | Green (VIC) | 35 | 7.8 | 2.9 | 16.4 | 36.4 | 37 |
|  | Yellow (NED) | 23 | 10.5 | 3.4 | 20.7 | 44.3 | 45 |
|  | Red (PET) | 37 | 8.0 | 3.0 | 17.0 | 37.0 | 37 |
| 310 | Blue (6-FAM) | 37 | 6.5 | 3.0 | 15.4 | 36.2 | 37 |
|  | Green (VIC) | 45 | 7.5 | 3.5 | 18.0 | 42.7 | 43 |
|  | Yellow (NED) | 25 | 8.9 | 2.9 | 17.7 | 38.3 | 39 |
|  | Red (PET) | 49 | 9.1 | 3.5 | 19.8 | 44.6 | 45 |
| 310 | Blue (6-FAM) | 28 | 7.0 | 2.9 | 15.8 | 36.5 | 37 |
|  | Green (VIC) | 41 | 9.0 | 4.2 | 21.7 | 51.3 | 52 |
|  | Yellow (NED) | 34 | 11.0 | 4.4 | 24.2 | 54.9 | 55 |
|  | Red (PET) | 55 | 11.5 | 5.1 | 26.9 | 62.9 | 63 |
| 3130 | Blue (6-FAM) | 26 | 6.5 | 3.2 | 16.0 | 38.1 | 39 |
|  | Green (VIC) | 59 | 6.3 | 3.7 | 17.3 | 42.9 | 43 |
|  | Yellow (NED) | 72 | 7.3 | 3.8 | 18.8 | 45.7 | 46 |
|  | Red (PET) | 51 | 9.0 | 3.5 | 19.4 | 43.7 | 44 |
| 3130 | Blue (6-FAM) | 26 | 6.4 | 3.3 | 16.2 | 39.1 | 40 |
|  | Green (VIC) | 37 | 7.0 | 5.2 | 22.5 | 58.7 | 59 |
|  | Yellow (NED) | 37 | 8.0 | 3.3 | 17.8 | 40.7 | 41 |
|  | Red (PET) | 61 | 8.7 | 3.8 | 20.1 | 46.8 | 47 |
| 3130 | Blue (6-FAM) | 45 | 7.0 | 5.9 | 24.8 | 66.1 | 67 |
|  | Green (VIC) | 37 | 6.8 | 5.0 | 21.8 | 56.6 | 57 |
|  | Yellow (NED) | 54 | 8,0 | 3.6 | 18.8 | 43.9 | 44 |
|  | Red (PET) | 51 | 8.7 | 3.7 | 19.8 | 45.7 | 46 |
| 3500 | Blue (6-FAM) | 34 | 7.3 | 4.4 | 20.4 | 51.1 | 52 |
|  | Green (VIC) | 42 | 7.7 | 3.1 | 17.0 | 38.5 | 39 |
|  | Yellow (NED) | 37 | 15.1 | 5.1 | 30.3 | 65.9 | 66 |
|  | Red (PET) | 48 | 22.6 | 8.1 | 46.8 | 103.3 | 104 |
| 3500 | Blue (6-FAM) | 26 | 6.7 | 2.7 | 14.8 | 33.8 | 34 |
|  | Green (VIC) | 42 | 7.8 | 3.6 | 18.6 | 43.9 | 44 |
|  | Yellow (NED) | 46 | 15.2 | 5.5 | 31.6 | 69.8 | 70 |
|  | Red (PET) | 54 | 22.4 | 8.1 | 46.7 | 103.4 | 104 |
| 3500 | Blue (6-FAM) | 41 | 7.5 | 5.4 | 23.7 | 61.6 | 62 |
|  | Green (VIC) | 36 | 7.5 | 3.2 | 17.2 | 39.7 | 40 |
|  | Yellow (NED) | 51 | 12.6 | 5.1 | 27.8 | 63.3 | 64 |
|  | Red (PET) | 62 | 18.9 | 7.7 | 42.0 | 95.9 | 96 |
| 3500 | Blue (6-FAM) | 43 | 6.8 | 3.8 | 18.2 | 44.8 | 45 |
|  | Green (VIC) | 62 | 8.8 | 4.0 | 20.9 | 49.2 | 50 |
|  | Yellow (NED) | 52 | 16.4 | 6.2 | 25.1 | 78.6 | 79 |
|  | Red (PET) | 65 | 23.4 | 9.1 | 50.7 | 114.2 | 115 |
| 3500 | Blue (6-FAM) | 35 | 7.0 | 2.3 | 14.0 | 30.3 | 31 |
|  | Green (VIC) | 42 | 7.6 | 2.5 | 15.2 | 32.9 | 33 |
|  | Yellow (NED) | 43 | 14.5 | 4.9 | 29.1 | 63.4 | 64 |
|  | Red (PET) | 54 | 20.6 | 7.3 | 42.4 | 93.3 | 94 |
| 3500 | Blue (6-FAM) | 23 | 6.6 | 2.6 | 14.3 | 32.2 | 33 |
|  | Green (VIC) | 56 | 8.3 | 5.3 | 24.1 | 61.1 | 62 |
|  | Yellow (NED) | 40 | 15.1 | 5.2 | 30.8 | 67.5 | 68 |
|  | Red (PET) | 54 | 22.5 | 8.4 | 47.7 | 106.5 | 107 |
| 3500 | Blue (6-FAM) | 37 | 6.2 | 2.6 | 14.0 | 32,0 | 32 |
|  | Green (VIC) | 59 | 8.4 | 4.6 | 22.1 | 54.0 | 54 |
|  | Yellow (NED) | 48 | 14.6 | 5.5 | 31.0 | 69.2 | 70 |
|  | Red (PET) | 48 | 21.7 | 7.6 | 44.5 | 97.8 | 98 |
| 3500 | Blue (6-FAM) | 30 | 9.3 | 6.3 | 28.3 | 72.8 | 73 |
|  | Green (VIC) | 42 | 8.1 | 3.4 | 18.3 | 42.1 | 43 |
|  | Yellow (NED) | 43 | 14.8 | 5.3 | 30.8 | 68.2 | 69 |
|  | Red (PET) | 90 | 24.1 | 9.2 | 51.6 | 115.8 | 116 |
| 3500XL | Blue (6-FAM) | 29 | 7.4 | 4.5 | 20.8 | 52.1 | 53 |
|  | Green (VIC) | 51 | 8.6 | 4.3 | 21.5 | 51.6 | 52 |
|  | Yellow (NED) | 45 | 16.5 | 5.5 | 33.2 | 72.0 | 72 |
|  | Red (PET) | 73 | 22.9 | 8.5 | 48.4 | 107.9 | 108 |
| 3500XL | Blue (6-FAM) | 34 | 6.4 | 3.8 | 19.0 | 45.5 | 46 |
|  | Green (VIC) | 50 | 7.6 | 3.8 | 19.0 | 45.5 | 46 |
|  | Yellow (NED) | 42 | 14.8 | 5.8 | 32.1 | 72.4 | 73 |
|  | Red (PET) | 61 | 21.1 | 9.1 | 48.4 | 112.3 | 113 |

| 3730 | Blue (6-FAM) | 26 | 6.3 | 2.2 | 12.9 | 28.2 | 29 |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Green (VIC) | 42 | 7.4 | 2.6 | 15.4 | 33.9 | 34 |
|  | Yellow (NED) | 40 | 14.6 | 4.8 | 29.1 | 62.9 | 63 |
|  | Red (PET) | 54 | 22.4 | 7.7 | 45.4 | 99.1 | 100 |
| SeqStudio | Blue (6-FAM) | 22 | 4.8 | 2.1 | 11.2 | 26.3 | 27 |
|  | Green (VIC) | 29 | 5.1 | 2.2 | 11.7 | 27.2 | 28 |
|  | Yellow (NED) | 22 | 10.6 | 3.3 | 20.6 | 43.9 | 44 |
|  | Red (PET) | 56 | 21.1 | 7.0 | 42.1 | 91.2 | 92 |
| SeqStudio | Blue (6-FAM) | 53 | 8.3 | 3.5 | 18.9 | 43.5 | 44 |
|  | Green (VIC) | 40 | 8.3 | 3.3 | 18.3 | 41.7 | 42 |
|  | Yellow (NED) | 60 | 16.7 | 5.2 | 32.5 | 69.2 | 70 |
|  | Red (PET) | 70 | 40.9 | 11.9 | 76.7 | 160.2 | 161 |

Limit of Detection (LOD) is the average plus 3 SD, the RFU value which 99.7% of the background noise peak should be observed; Limit of Quantification (LOQ) is the average plus 10 SD,; provides an upper limit value below which all or nearly all background “noise” would be expected to fall (99.87%  of the background noise peak); Analytical Threshold (AT) is obtained rounding the LOQ of each dye channel to the nearest superior integer; Maximum Peak Height, in RFU, is the maximum peak height found on negative triplicate control in that channel; Average Peak Height, in RFU,  is the mean of negative control peaks found in that channel.